# CS650 Frequency-Domain Reflectometer (FDR) Sensor

The CS650 Water Content Reflectometer measures volumetric water content (VWC) by probing the soil’s dielectric constant via high-frequency electromagnetic waves. Its key specifications and operating principles are:

- Measurement Principle: Frequency-Domain Reflectometry (FDR) at a nominal frequency of 70 MHz. A burst of RF energy is emitted along the waveguide rods; the reflected signal phase shift is proportional to the soil’s dielectric permittivity ε.

- Probe Geometry & Volume: Three stainless-steel waveguides, 10 cm long and 2.5 mm diameter, form an equilateral triangle. The effective sampling volume is approximately a cylinder of length 10 cm and radius 3–5 cm around the probes.

- Calibration & Conversion: The raw phase shift ϕ is processed inside the CR6 datalogger to compute ε, then converted to VWC (θ) using a factory-tuned calibration curve:

θ=a0+a1 ε+a2 ε2 ,

- where a0,a1,a2 are sensor-specific coefficients.

- Accuracy & Resolution:

  - Typical accuracy: ±0.02 (±2

  - Resolution: 0.0001 m³/m³

  - Response time: ≪1 s

- Environmental Range: θ∈[0, 0.5] m³/m³; soil temperature range −40 °C to +60 °C.

- Data Logging: Paired with a Campbell Scientific CR6 datalogger, which timestamps and stores measurements at user-defined intervals, and can apply temperature and bulk density corrections in real time.